## SSC style Output Cert
Flow Lifecycle Services |  | Certificate of Lowflow Air Calibration | [6 empty] | 14512664
Product Code | Serial ID |  | Order ID |  |  |  | Customer Tag |  | Customer
CMF010H323N2BAEZZX_26289 | 14512664 |  | 10214866 | [5 empty] | Alcoa Inc
PUCK800 | 25843369
FLOW VERIFICATION - FLOW CAL FACTORS SAME AS INITIAL CAL (AS FOUND / AS LEFT)
Process |  |  | Detail
Customer PO: | 140506627
Process Date: | 2016-09-09
Reference Meter: | CMFS007; SN. 12095647
Stand Uncertainty: | ± 1%
Test Gas: | Argon
Freq: | Meter Output, Hertz (pulses/sec)
Press: | Meter INLET static pressure (100 psia)
Temp: | 72 F
Mdot: | Mass flowrate in kilograms per minute
Batch Time: | 120 sec
Standard Dev (%) | 0.086215
Results
Status: | PASS
D1: | 0
D2: | 1
K1: | 10120.3398
K2: | 11090.71
DT: | 2.82
FD: | 261.0988 |  | Number | Flow Rate | Meter | Reference | Error | Specification | Specification
DTG: | 0 |  |  | (Kg/min) | Total (Kg) | Total (Kg) | (%) | (+%) | (-%)
DFQ1: | 0 |  | 1 | 0.0014 | 0.002912 | 0.002905 | 0.24655 | 1 | -1
DFQ2: | 0 |  | 2 | 0.0014 | 0.002885 | 0.002894 | 0.29732 | 1 | -1
FlowCal: | .442603.41 |  | 3 | 0.0014 | 0.002945 | 0.002937 | 0.28126 | 1 | -1
FFQ: | 0 |  | 4 | 0.005 | 0.01016 | 0.010142 | 0.17922 | 1 | -1
FTG: | 0 |  | 5 | 0.005 | 0.010175 | 0.010167 | 0.071473 | 1 | -1
DensCal: | 10120110912.82 |  | 6 | 0.005 | 0.010199 | 0.010183 | 0.15395 | 1 | -1
FCF: | 0.4426
FT: | 3.41
Rabi Gurung - Repair Support Eng.
Meter total is based on digital output. |  |  | Custom low flow calibration performed in the Repair Lab

## Information
TEST DATE & TIME | 10:54:38 09.09.16
TEST COMMENTS:
SENSOR DATA
Serial Number: | 14512664
Model:
Size:
Material: | Special
Press Rating:
Temp Rating:
SENSOR CALIBRATION
Flow Calibration Factor: | 0.4426
D1: | 0
D2: | 1
K1: | 10120.3398
K2: | 11090.71
FD: | 261.0988
Density Temp Coeff.: | 2.82
TRANSMITTER DATA
Serial Number:
Model: | 660
Mass Flow Units: | Grams/minute
Density Units: | Grams/cubic centimeter
Temperature Units: | Degrees Fahrenheit
Mass Flow Cutoff: | 0
Mass Flow Damping: | 0.8
Flow Direction: | Negate/Forward only
Frequency Flow Rate: | 591.8917
Frequency Span: | 277.2352
Meter Specifications
Flow Accuracy: |  | %
Flow Repeatability: |  | %
Zero Stability: |  | lb/min
Density Accuracy: |  | g/cc
Density Repeatability: |  | g/cc
Nominal Flow Rate |  | lb/min
Meter zero | 0.029547 | nsec
Customer data
Status | As found | as found/as left
Customer name
Customer address
Sales order number
Technician name
ref meter
Product code
Fluid | H2O
Fluid temp

## RawData
Data Point | UUT Flow | UUT Temp | UUT Density | UUT Zero nsec | Set Point | UUT Dig Total1 | Ref | Error Dig-Scale % | Start Scale Corr lb | End Scale Gross lb | End Scale Corr lb | Real-time Bouyancy | Inlet Press psig | Diff Press psid | Inlet Temp °C | Exit Temp °C | Fluid Density g/cc | Ambient Press psia | Ambient Temp °C | Ambient % RH | VFD Speed % | Batch Time sec | Rdgs Avg'd | Timestamp | Analog Output mA
1 |  |  |  |  | 0.0014 | 0.002912 | 0.002905 | 0.24655 | [17 empty] | 0.26333 | 0.29246
2 |  |  |  |  | 0.0014 | 0.002885 | 0.002894 | 0.29732 | [17 empty] | 0.26213 | 0.29601
3 |  |  |  |  | 0.0014 | 0.002945 | 0.002937 | 0.28126 | [17 empty] | 0.26652 | 0.29228
4 |  |  |  |  | 0.0014 | 0.00289 | 0.002883 | 0.24272 | [17 empty] | 0.26328 | 0.28874
5 |  |  |  |  | 0.005 | 0.01016 | 0.010142 | 0.17922 | [17 empty] | 0.93834 | 1.026
6 |  |  |  |  | 0.005 | 0.010175 | 0.010167 | 0.071473 | [17 empty] | 0.94222 | 1.0599
7 |  |  |  |  | 0.005 | 0.010199 | 0.010183 | 0.15395 | [17 empty] | 0.9423 | 1.0347
8 |  |  |  |  | 0.005 | 0.010191 | 0.010182 | 0.084591 | [17 empty] | 0.93662 | 1.0168
9 |  |  |  |  | 0.0014 | 0.010214 | 0.010211 | 0.030655 | [17 empty] | 0.94195 | 1.0322
10 |  |  |  |  | 0.0014 | 0.010232 | 0.010234 | 0.020061 | [17 empty] | 0.94173 | 1.0565
11 |  |  |  |  | 0.0014 | 0.010231 | 0.010233 | 0.014644 | [17 empty] | 0.93928 | 1.0503

## Error Limits
Error | Flow |  | FALSE
Limit | Rate
Turndown | lb/min | % Flow | + Limit | - Limit | + D Limit | - D Limit
1 | 591.8917 | #DIV/0! | 0 | 0 | 0 | 0
1.5 | 394.5944 | #DIV/0! | 0 | 0 | 0 | 0
2 | 295.9458 | #DIV/0! | 0 | 0 | 0 | 0
2.5 | 236.7567 | #DIV/0! | 0 | 0 | 0 | 0
3 | 197.2972 | #DIV/0! | 0 | 0 | 0 | 0
3.5 | 169.1119 | #DIV/0! | 0 | 0 | 0 | 0
4 | 147.9729 | #DIV/0! | 0 | 0 | 0 | 0
4.5 | 131.5315 | #DIV/0! | 0 | 0 | 0 | 0
5 | 118.3783 | #DIV/0! | 0 | 0 | 0 | 0
5.5 | 107.6167 | #DIV/0! | 0 | 0 | 0 | 0
6 | 98.6486 | #DIV/0! | 0 | 0 | 0 | 0
6.5 | 91.0603 | #DIV/0! | 0 | 0 | 0 | 0
7 | 84.556 | #DIV/0! | 0 | 0 | 0 | 0
7.5 | 78.9189 | #DIV/0! | 0 | 0 | 0 | 0
8 | 73.9865 | #DIV/0! | 0 | 0 | 0 | 0
8.5 | 69.6343 | #DIV/0! | 0 | 0 | 0 | 0
9 | 65.7657 | #DIV/0! | 0 | 0 | 0 | 0
9.5 | 62.3044 | #DIV/0! | 0 | 0 | 0 | 0
10 | 59.1892 | #DIV/0! | 0 | 0 | 0 | 0
10.5 | 56.3706 | #DIV/0! | 0 | 0 | 0 | 0
11 | 53.8083 | #DIV/0! | 0 | 0 | 0 | 0
11.5 | 51.4688 | #DIV/0! | 0 | 0 | 0 | 0
12 | 49.3243 | #DIV/0! | 0 | 0 | 0 | 0
12.5 | 47.3513 | #DIV/0! | 0 | 0 | 0 | 0
13 | 45.5301 | #DIV/0! | 0 | 0 | 0 | 0
13.5 | 43.8438 | #DIV/0! | 0 | 0 | 0 | 0
14 | 42.278 | #DIV/0! | 0 | 0 | 0 | 0
14.5 | 40.8201 | #DIV/0! | 0 | 0 | 0 | 0
15 | 39.4594 | #DIV/0! | 0 | 0 | 0 | 0
15.5 | 38.1866 | #DIV/0! | 0 | 0 | 0 | 0
16 | 36.9932 | #DIV/0! | 0 | 0 | 0 | 0
16.5 | 35.8722 | #DIV/0! | 0 | 0 | 0 | 0
17 | 34.8172 | #DIV/0! | 0 | 0 | 0 | 0
17.5 | 33.8224 | #DIV/0! | 0 | 0 | 0 | 0
18 | 32.8829 | #DIV/0! | 0 | 0 | 0 | 0
18.5 | 31.9941 | #DIV/0! | 0 | 0 | 0 | 0
19 | 31.1522 | #DIV/0! | 0 | 0 | 0 | 0
19.5 | 30.3534 | #DIV/0! | 0 | 0 | 0 | 0
20 | 29.5946 | #DIV/0! | 0 | 0 | 0 | 0
20.5 | 28.8728 | #DIV/0! | 0 | 0 | 0 | 0
21 | 28.1853 | #DIV/0! | 0 | 0 | 0 | 0
21.5 | 27.5298 | #DIV/0! | 0 | 0 | 0 | 0
22 | 26.9042 | #DIV/0! | 0 | 0 | 0 | 0
22.5 | 26.3063 | #DIV/0! | 0 | 0 | 0 | 0
23 | 25.7344 | #DIV/0! | 0 | 0 | 0 | 0
23.5 | 25.1869 | #DIV/0! | 0 | 0 | 0 | 0
24 | 24.6622 | #DIV/0! | 0 | 0 | 0 | 0
24.5 | 24.1588 | #DIV/0! | 0 | 0 | 0 | 0
25 | 23.6757 | #DIV/0! | 0 | 0 | 0 | 0
25.5 | 23.2114 | #DIV/0! | 0 | 0 | 0 | 0
26 | 22.7651 | #DIV/0! | 0 | 0 | 0 | 0
26.5 | 22.3355 | #DIV/0! | 0 | 0 | 0 | 0
27 | 21.9219 | #DIV/0! | 0 | 0 | 0 | 0
27.5 | 21.5233 | #DIV/0! | 0 | 0 | 0 | 0
28 | 21.139 | #DIV/0! | 0 | 0 | 0 | 0
28.5 | 20.7681 | #DIV/0! | 0 | 0 | 0 | 0
29 | 20.4101 | #DIV/0! | 0 | 0 | 0 | 0
29.5 | 20.0641 | #DIV/0! | 0 | 0 | 0 | 0
30 | 19.7297 | #DIV/0! | 0 | 0 | 0 | 0
30.5 | 19.4063 | #DIV/0! | 0 | 0 | 0 | 0
31 | 19.0933 | #DIV/0! | 0 | 0 | 0 | 0
31.5 | 18.7902 | #DIV/0! | 0 | 0 | 0 | 0
32 | 18.4966 | #DIV/0! | 0 | 0 | 0 | 0
32.5 | 18.2121 | #DIV/0! | 0 | 0 | 0 | 0
33 | 17.9361 | #DIV/0! | 0 | 0 | 0 | 0
33.5 | 17.6684 | #DIV/0! | 0 | 0 | 0 | 0
34 | 17.4086 | #DIV/0! | 0 | 0 | 0 | 0
34.5 | 17.1563 | #DIV/0! | 0 | 0 | 0 | 0
35 | 16.9112 | #DIV/0! | 0 | 0 | 0 | 0
35.5 | 16.673 | #DIV/0! | 0 | 0 | 0 | 0
36 | 16.4414 | #DIV/0! | 0 | 0 | 0 | 0
36.5 | 16.2162 | #DIV/0! | 0 | 0 | 0 | 0
37 | 15.9971 | #DIV/0! | 0 | 0 | 0 | 0
37.5 | 15.7838 | #DIV/0! | 0 | 0 | 0 | 0
38 | 15.5761 | #DIV/0! | 0 | 0 | 0 | 0
38.5 | 15.3738 | #DIV/0! | 0 | 0 | 0 | 0
39 | 15.1767 | #DIV/0! | 0 | 0 | 0 | 0
39.5 | 14.9846 | #DIV/0! | 0 | 0 | 0 | 0
40 | 14.7973 | #DIV/0! | 0 | 0 | 0 | 0
40.5 | 14.6146 | #DIV/0! | 0 | 0 | 0 | 0
41 | 14.4364 | #DIV/0! | 0 | 0 | 0 | 0
41.5 | 14.2624 | #DIV/0! | 0 | 0 | 0 | 0
42 | 14.0927 | #DIV/0! | 0 | 0 | 0 | 0
42.5 | 13.9269 | #DIV/0! | 0 | 0 | 0 | 0
43 | 13.7649 | #DIV/0! | 0 | 0 | 0 | 0
43.5 | 13.6067 | #DIV/0! | 0 | 0 | 0 | 0
44 | 13.4521 | #DIV/0! | 0 | 0 | 0 | 0
44.5 | 13.3009 | #DIV/0! | 0 | 0 | 0 | 0
45 | 13.1531 | #DIV/0! | 0 | 0 | 0 | 0
45.5 | 13.0086 | #DIV/0! | 0 | 0 | 0 | 0
46 | 12.8672 | #DIV/0! | 0 | 0 | 0 | 0
46.5 | 12.7289 | #DIV/0! | 0 | 0 | 0 | 0
47 | 12.5934 | #DIV/0! | 0 | 0 | 0 | 0
47.5 | 12.4609 | #DIV/0! | 0 | 0 | 0 | 0
48 | 12.3311 | #DIV/0! | 0 | 0 | 0 | 0
48.5 | 12.204 | #DIV/0! | 0 | 0 | 0 | 0
49 | 12.0794 | #DIV/0! | 0 | 0 | 0 | 0
49.5 | 11.9574 | #DIV/0! | 0 | 0 | 0 | 0
50 | 11.8378 | #DIV/0! | 0 | 0 | 0 | 0
50.5 | 11.7206 | #DIV/0! | 0 | 0 | 0 | 0
51 | 11.6057 | #DIV/0! | 0 | 0 | 0 | 0
51.5 | 11.493 | #DIV/0! | 0 | 0 | 0 | 0
52 | 11.3825 | #DIV/0! | 0 | 0 | 0 | 0
52.5 | 11.2741 | #DIV/0! | 0 | 0 | 0 | 0
53 | 11.1678 | #DIV/0! | 0 | 0 | 0 | 0
53.5 | 11.0634 | #DIV/0! | 0 | 0 | 0 | 0
54 | 10.961 | #DIV/0! | 0 | 0 | 0 | 0
54.5 | 10.8604 | #DIV/0! | 0 | 0 | 0 | 0
55 | 10.7617 | #DIV/0! | 0 | 0 | 0 | 0
55.5 | 10.6647 | #DIV/0! | 0 | 0 | 0 | 0
56 | 10.5695 | #DIV/0! | 0 | 0 | 0 | 0
56.5 | 10.476 | #DIV/0! | 0 | 0 | 0 | 0
57 | 10.3841 | #DIV/0! | 0 | 0 | 0 | 0
57.5 | 10.2938 | #DIV/0! | 0 | 0 | 0 | 0
58 | 10.205 | #DIV/0! | 0 | 0 | 0 | 0
58.5 | 10.1178 | #DIV/0! | 0 | 0 | 0 | 0
59 | 10.0321 | #DIV/0! | 0 | 0 | 0 | 0
59.5 | 9.9478 | #DIV/0! | 0 | 0 | 0 | 0
60 | 9.8649 | #DIV/0! | 0 | 0 | 0 | 0
60.5 | 9.7833 | #DIV/0! | 0 | 0 | 0 | 0
61 | 9.7031 | #DIV/0! | 0 | 0 | 0 | 0
61.5 | 9.6243 | #DIV/0! | 0 | 0 | 0 | 0
62 | 9.5466 | #DIV/0! | 0 | 0 | 0 | 0
62.5 | 9.4703 | #DIV/0! | 0 | 0 | 0 | 0
63 | 9.3951 | #DIV/0! | 0 | 0 | 0 | 0
63.5 | 9.3211 | #DIV/0! | 0 | 0 | 0 | 0
64 | 9.2483 | #DIV/0! | 0 | 0 | 0 | 0
64.5 | 9.1766 | #DIV/0! | 0 | 0 | 0 | 0
65 | 9.106 | #DIV/0! | 0 | 0 | 0 | 0
65.5 | 9.0365 | #DIV/0! | 0 | 0 | 0 | 0
66 | 8.9681 | #DIV/0! | 0 | 0 | 0 | 0
66.5 | 8.9006 | #DIV/0! | 0 | 0 | 0 | 0
67 | 8.8342 | #DIV/0! | 0 | 0 | 0 | 0
67.5 | 8.7688 | #DIV/0! | 0 | 0 | 0 | 0
68 | 8.7043 | #DIV/0! | 0 | 0 | 0 | 0
68.5 | 8.6408 | #DIV/0! | 0 | 0 | 0 | 0
69 | 8.5781 | #DIV/0! | 0 | 0 | 0 | 0
69.5 | 8.5164 | #DIV/0! | 0 | 0 | 0 | 0
70 | 8.4556 | #DIV/0! | 0 | 0 | 0 | 0
70.5 | 8.3956 | #DIV/0! | 0 | 0 | 0 | 0
71 | 8.3365 | #DIV/0! | 0 | 0 | 0 | 0
71.5 | 8.2782 | #DIV/0! | 0 | 0 | 0 | 0
72 | 8.2207 | #DIV/0! | 0 | 0 | 0 | 0
72.5 | 8.164 | #DIV/0! | 0 | 0 | 0 | 0
73 | 8.1081 | #DIV/0! | 0 | 0 | 0 | 0
73.5 | 8.0529 | #DIV/0! | 0 | 0 | 0 | 0
74 | 7.9985 | #DIV/0! | 0 | 0 | 0 | 0
74.5 | 7.9449 | #DIV/0! | 0 | 0 | 0 | 0
75 | 7.8919 | #DIV/0! | 0 | 0 | 0 | 0
75.5 | 7.8396 | #DIV/0! | 0 | 0 | 0 | 0
76 | 7.788 | #DIV/0! | 0 | 0 | 0 | 0
76.5 | 7.7371 | #DIV/0! | 0 | 0 | 0 | 0
77 | 7.6869 | #DIV/0! | 0 | 0 | 0 | 0
77.5 | 7.6373 | #DIV/0! | 0 | 0 | 0 | 0
78 | 7.5884 | #DIV/0! | 0 | 0 | 0 | 0
78.5 | 7.54 | #DIV/0! | 0 | 0 | 0 | 0
79 | 7.4923 | #DIV/0! | 0 | 0 | 0 | 0
79.5 | 7.4452 | #DIV/0! | 0 | 0 | 0 | 0
80 | 7.3986 | #DIV/0! | 0 | 0 | 0 | 0
80.5 | 7.3527 | #DIV/0! | 0 | 0 | 0 | 0
81 | 7.3073 | #DIV/0! | 0 | 0 | 0 | 0
81.5 | 7.2625 | #DIV/0! | 0 | 0 | 0 | 0
82 | 7.2182 | #DIV/0! | 0 | 0 | 0 | 0
82.5 | 7.1744 | #DIV/0! | 0 | 0 | 0 | 0
83 | 7.1312 | #DIV/0! | 0 | 0 | 0 | 0
83.5 | 7.0885 | #DIV/0! | 0 | 0 | 0 | 0
84 | 7.0463 | #DIV/0! | 0 | 0 | 0 | 0
84.5 | 7.0046 | #DIV/0! | 0 | 0 | 0 | 0
85 | 6.9634 | #DIV/0! | 0 | 0 | 0 | 0
85.5 | 6.9227 | #DIV/0! | 0 | 0 | 0 | 0
86 | 6.8825 | #DIV/0! | 0 | 0 | 0 | 0
86.5 | 6.8427 | #DIV/0! | 0 | 0 | 0 | 0
87 | 6.8034 | #DIV/0! | 0 | 0 | 0 | 0
87.5 | 6.7645 | #DIV/0! | 0 | 0 | 0 | 0
88 | 6.726 | #DIV/0! | 0 | 0 | 0 | 0
88.5 | 6.688 | #DIV/0! | 0 | 0 | 0 | 0
89 | 6.6505 | #DIV/0! | 0 | 0 | 0 | 0
89.5 | 6.6133 | #DIV/0! | 0 | 0 | 0 | 0
90 | 6.5766 | #DIV/0! | 0 | 0 | 0 | 0
90.5 | 6.5402 | #DIV/0! | 0 | 0 | 0 | 0
91 | 6.5043 | #DIV/0! | 0 | 0 | 0 | 0
91.5 | 6.4688 | #DIV/0! | 0 | 0 | 0 | 0
92 | 6.4336 | #DIV/0! | 0 | 0 | 0 | 0
92.5 | 6.3988 | #DIV/0! | 0 | 0 | 0 | 0
93 | 6.3644 | #DIV/0! | 0 | 0 | 0 | 0
93.5 | 6.3304 | #DIV/0! | 0 | 0 | 0 | 0
94 | 6.2967 | #DIV/0! | 0 | 0 | 0 | 0
94.5 | 6.2634 | #DIV/0! | 0 | 0 | 0 | 0
95 | 6.2304 | #DIV/0! | 0 | 0 | 0 | 0
95.5 | 6.1978 | #DIV/0! | 0 | 0 | 0 | 0
96 | 6.1655 | #DIV/0! | 0 | 0 | 0 | 0
96.5 | 6.1336 | #DIV/0! | 0 | 0 | 0 | 0
97 | 6.102 | #DIV/0! | 0 | 0 | 0 | 0
97.5 | 6.0707 | #DIV/0! | 0 | 0 | 0 | 0
98 | 6.0397 | #DIV/0! | 0 | 0 | 0 | 0
98.5 | 6.0091 | #DIV/0! | 0 | 0 | 0 | 0
99 | 5.9787 | #DIV/0! | 0 | 0 | 0 | 0
99.5 | 5.9487 | #DIV/0! | 0 | 0 | 0 | 0
100 | 5.9189 | #DIV/0! | 0 | 0 | 0 | 0
100.5 | 5.8895 | #DIV/0! | 0 | 0 | 0 | 0
101 | 5.8603 | #DIV/0! | 0 | 0 | 0 | 0
101.5 | 5.8314 | #DIV/0! | 0 | 0 | 0 | 0
102 | 5.8029 | #DIV/0! | 0 | 0 | 0 | 0
102.5 | 5.7746 | #DIV/0! | 0 | 0 | 0 | 0
103 | 5.7465 | #DIV/0! | 0 | 0 | 0 | 0
103.5 | 5.7188 | #DIV/0! | 0 | 0 | 0 | 0
104 | 5.6913 | #DIV/0! | 0 | 0 | 0 | 0
104.5 | 5.664 | #DIV/0! | 0 | 0 | 0 | 0
105 | 5.6371 | #DIV/0! | 0 | 0 | 0 | 0
105.5 | 5.6103 | #DIV/0! | 0 | 0 | 0 | 0
106 | 5.5839 | #DIV/0! | 0 | 0 | 0 | 0
106.5 | 5.5577 | #DIV/0! | 0 | 0 | 0 | 0
107 | 5.5317 | #DIV/0! | 0 | 0 | 0 | 0
107.5 | 5.506 | #DIV/0! | 0 | 0 | 0 | 0
108 | 5.4805 | #DIV/0! | 0 | 0 | 0 | 0
108.5 | 5.4552 | #DIV/0! | 0 | 0 | 0 | 0
109 | 5.4302 | #DIV/0! | 0 | 0 | 0 | 0
109.5 | 5.4054 | #DIV/0! | 0 | 0 | 0 | 0
110 | 5.3808 | #DIV/0! | 0 | 0 | 0 | 0
110.5 | 5.3565 | #DIV/0! | 0 | 0 | 0 | 0
111 | 5.3324 | #DIV/0! | 0 | 0 | 0 | 0
111.5 | 5.3084 | #DIV/0! | 0 | 0 | 0 | 0
112 | 5.2847 | #DIV/0! | 0 | 0 | 0 | 0
112.5 | 5.2613 | #DIV/0! | 0 | 0 | 0 | 0
113 | 5.238 | #DIV/0! | 0 | 0 | 0 | 0
113.5 | 5.2149 | #DIV/0! | 0 | 0 | 0 | 0
114 | 5.192 | #DIV/0! | 0 | 0 | 0 | 0
114.5 | 5.1694 | #DIV/0! | 0 | 0 | 0 | 0
115 | 5.1469 | #DIV/0! | 0 | 0 | 0 | 0
115.5 | 5.1246 | #DIV/0! | 0 | 0 | 0 | 0
116 | 5.1025 | #DIV/0! | 0 | 0 | 0 | 0
116.5 | 5.0806 | #DIV/0! | 0 | 0 | 0 | 0
117 | 5.0589 | #DIV/0! | 0 | 0 | 0 | 0
117.5 | 5.0374 | #DIV/0! | 0 | 0 | 0 | 0
118 | 5.016 | #DIV/0! | 0 | 0 | 0 | 0
118.5 | 4.9949 | #DIV/0! | 0 | 0 | 0 | 0
119 | 4.9739 | #DIV/0! | 0 | 0 | 0 | 0
119.5 | 4.9531 | #DIV/0! | 0 | 0 | 0 | 0
120 | 4.9324 | #DIV/0! | 0 | 0 | 0 | 0
120.5 | 4.912 | #DIV/0! | 0 | 0 | 0 | 0
121 | 4.8917 | #DIV/0! | 0 | 0 | 0 | 0
121.5 | 4.8715 | #DIV/0! | 0 | 0 | 0 | 0
122 | 4.8516 | #DIV/0! | 0 | 0 | 0 | 0
122.5 | 4.8318 | #DIV/0! | 0 | 0 | 0 | 0
123 | 4.8121 | #DIV/0! | 0 | 0 | 0 | 0
123.5 | 4.7926 | #DIV/0! | 0 | 0 | 0 | 0
124 | 4.7733 | #DIV/0! | 0 | 0 | 0 | 0
124.5 | 4.7541 | #DIV/0! | 0 | 0 | 0 | 0
125 | 4.7351 | #DIV/0! | 0 | 0 | 0 | 0
125.5 | 4.7163 | #DIV/0! | 0 | 0 | 0 | 0
126 | 4.6976 | #DIV/0! | 0 | 0 | 0 | 0
126.5 | 4.679 | #DIV/0! | 0 | 0 | 0 | 0
127 | 4.6606 | #DIV/0! | 0 | 0 | 0 | 0
127.5 | 4.6423 | #DIV/0! | 0 | 0 | 0 | 0
128 | 4.6242 | #DIV/0! | 0 | 0 | 0 | 0
128.5 | 4.6062 | #DIV/0! | 0 | 0 | 0 | 0
129 | 4.5883 | #DIV/0! | 0 | 0 | 0 | 0
129.5 | 4.5706 | #DIV/0! | 0 | 0 | 0 | 0
130 | 4.553 | #DIV/0! | 0 | 0 | 0 | 0
130.5 | 4.5356 | #DIV/0! | 0 | 0 | 0 | 0
131 | 4.5183 | #DIV/0! | 0 | 0 | 0 | 0
131.5 | 4.5011 | #DIV/0! | 0 | 0 | 0 | 0
132 | 4.484 | #DIV/0! | 0 | 0 | 0 | 0
132.5 | 4.4671 | #DIV/0! | 0 | 0 | 0 | 0
133 | 4.4503 | #DIV/0! | 0 | 0 | 0 | 0
133.5 | 4.4336 | #DIV/0! | 0 | 0 | 0 | 0
134 | 4.4171 | #DIV/0! | 0 | 0 | 0 | 0
134.5 | 4.4007 | #DIV/0! | 0 | 0 | 0 | 0
135 | 4.3844 | #DIV/0! | 0 | 0 | 0 | 0
135.5 | 4.3682 | #DIV/0! | 0 | 0 | 0 | 0
136 | 4.3521 | #DIV/0! | 0 | 0 | 0 | 0
136.5 | 4.3362 | #DIV/0! | 0 | 0 | 0 | 0
137 | 4.3204 | #DIV/0! | 0 | 0 | 0 | 0
137.5 | 4.3047 | #DIV/0! | 0 | 0 | 0 | 0
138 | 4.2891 | #DIV/0! | 0 | 0 | 0 | 0
138.5 | 4.2736 | #DIV/0! | 0 | 0 | 0 | 0
139 | 4.2582 | #DIV/0! | 0 | 0 | 0 | 0
139.5 | 4.243 | #DIV/0! | 0 | 0 | 0 | 0
140 | 4.2278 | #DIV/0! | 0 | 0 | 0 | 0
140.5 | 4.2128 | #DIV/0! | 0 | 0 | 0 | 0
141 | 4.1978 | #DIV/0! | 0 | 0 | 0 | 0
141.5 | 4.183 | #DIV/0! | 0 | 0 | 0 | 0
142 | 4.1683 | #DIV/0! | 0 | 0 | 0 | 0
142.5 | 4.1536 | #DIV/0! | 0 | 0 | 0 | 0
143 | 4.1391 | #DIV/0! | 0 | 0 | 0 | 0
143.5 | 4.1247 | #DIV/0! | 0 | 0 | 0 | 0
144 | 4.1104 | #DIV/0! | 0 | 0 | 0 | 0
144.5 | 4.0961 | #DIV/0! | 0 | 0 | 0 | 0
145 | 4.082 | #DIV/0! | 0 | 0 | 0 | 0
145.5 | 4.068 | #DIV/0! | 0 | 0 | 0 | 0
146 | 4.0541 | #DIV/0! | 0 | 0 | 0 | 0
146.5 | 4.0402 | #DIV/0! | 0 | 0 | 0 | 0
147 | 4.0265 | #DIV/0! | 0 | 0 | 0 | 0
147.5 | 4.0128 | #DIV/0! | 0 | 0 | 0 | 0
148 | 3.9993 | #DIV/0! | 0 | 0 | 0 | 0
148.5 | 3.9858 | #DIV/0! | 0 | 0 | 0 | 0
149 | 3.9724 | #DIV/0! | 0 | 0 | 0 | 0
149.5 | 3.9591 | #DIV/0! | 0 | 0 | 0 | 0
150 | 3.9459 | #DIV/0! | 0 | 0 | 0 | 0
150.5 | 3.9328 | #DIV/0! | 0 | 0 | 0 | 0
151 | 3.9198 | #DIV/0! | 0 | 0 | 0 | 0
151.5 | 3.9069 | #DIV/0! | 0 | 0 | 0 | 0
152 | 3.894 | #DIV/0! | 0 | 0 | 0 | 0
152.5 | 3.8813 | #DIV/0! | 0 | 0 | 0 | 0
153 | 3.8686 | #DIV/0! | 0 | 0 | 0 | 0
153.5 | 3.856 | #DIV/0! | 0 | 0 | 0 | 0
154 | 3.8435 | #DIV/0! | 0 | 0 | 0 | 0
154.5 | 3.831 | #DIV/0! | 0 | 0 | 0 | 0
155 | 3.8187 | #DIV/0! | 0 | 0 | 0 | 0
155.5 | 3.8064 | #DIV/0! | 0 | 0 | 0 | 0
156 | 3.7942 | #DIV/0! | 0 | 0 | 0 | 0
156.5 | 3.7821 | #DIV/0! | 0 | 0 | 0 | 0
157 | 3.77 | #DIV/0! | 0 | 0 | 0 | 0
157.5 | 3.758 | #DIV/0! | 0 | 0 | 0 | 0
158 | 3.7461 | #DIV/0! | 0 | 0 | 0 | 0
158.5 | 3.7343 | #DIV/0! | 0 | 0 | 0 | 0
159 | 3.7226 | #DIV/0! | 0 | 0 | 0 | 0
159.5 | 3.7109 | #DIV/0! | 0 | 0 | 0 | 0
160 | 3.6993 | #DIV/0! | 0 | 0 | 0 | 0
160.5 | 3.6878 | #DIV/0! | 0 | 0 | 0 | 0
161 | 3.6763 | #DIV/0! | 0 | 0 | 0 | 0
161.5 | 3.665 | #DIV/0! | 0 | 0 | 0 | 0
162 | 3.6537 | #DIV/0! | 0 | 0 | 0 | 0
162.5 | 3.6424 | #DIV/0! | 0 | 0 | 0 | 0
163 | 3.6312 | #DIV/0! | 0 | 0 | 0 | 0
163.5 | 3.6201 | #DIV/0! | 0 | 0 | 0 | 0
164 | 3.6091 | #DIV/0! | 0 | 0 | 0 | 0
164.5 | 3.5981 | #DIV/0! | 0 | 0 | 0 | 0
165 | 3.5872 | #DIV/0! | 0 | 0 | 0 | 0
165.5 | 3.5764 | #DIV/0! | 0 | 0 | 0 | 0
166 | 3.5656 | #DIV/0! | 0 | 0 | 0 | 0
166.5 | 3.5549 | #DIV/0! | 0 | 0 | 0 | 0
167 | 3.5443 | #DIV/0! | 0 | 0 | 0 | 0
167.5 | 3.5337 | #DIV/0! | 0 | 0 | 0 | 0
168 | 3.5232 | #DIV/0! | 0 | 0 | 0 | 0
168.5 | 3.5127 | #DIV/0! | 0 | 0 | 0 | 0
169 | 3.5023 | #DIV/0! | 0 | 0 | 0 | 0
169.5 | 3.492 | #DIV/0! | 0 | 0 | 0 | 0
170 | 3.4817 | #DIV/0! | 0 | 0 | 0 | 0
170.5 | 3.4715 | #DIV/0! | 0 | 0 | 0 | 0
171 | 3.4614 | #DIV/0! | 0 | 0 | 0 | 0
171.5 | 3.4513 | #DIV/0! | 0 | 0 | 0 | 0
172 | 3.4412 | #DIV/0! | 0 | 0 | 0 | 0
172.5 | 3.4313 | #DIV/0! | 0 | 0 | 0 | 0
173 | 3.4213 | #DIV/0! | 0 | 0 | 0 | 0
173.5 | 3.4115 | #DIV/0! | 0 | 0 | 0 | 0
174 | 3.4017 | #DIV/0! | 0 | 0 | 0 | 0
174.5 | 3.3919 | #DIV/0! | 0 | 0 | 0 | 0
175 | 3.3822 | #DIV/0! | 0 | 0 | 0 | 0
175.5 | 3.3726 | #DIV/0! | 0 | 0 | 0 | 0
176 | 3.363 | #DIV/0! | 0 | 0 | 0 | 0
176.5 | 3.3535 | #DIV/0! | 0 | 0 | 0 | 0
177 | 3.344 | #DIV/0! | 0 | 0 | 0 | 0
177.5 | 3.3346 | #DIV/0! | 0 | 0 | 0 | 0
178 | 3.3252 | #DIV/0! | 0 | 0 | 0 | 0
178.5 | 3.3159 | #DIV/0! | 0 | 0 | 0 | 0
179 | 3.3067 | #DIV/0! | 0 | 0 | 0 | 0
179.5 | 3.2974 | #DIV/0! | 0 | 0 | 0 | 0
180 | 3.2883 | #DIV/0! | 0 | 0 | 0 | 0
180.5 | 3.2792 | #DIV/0! | 0 | 0 | 0 | 0
181 | 3.2701 | #DIV/0! | 0 | 0 | 0 | 0
181.5 | 3.2611 | #DIV/0! | 0 | 0 | 0 | 0
182 | 3.2522 | #DIV/0! | 0 | 0 | 0 | 0
182.5 | 3.2432 | #DIV/0! | 0 | 0 | 0 | 0
183 | 3.2344 | #DIV/0! | 0 | 0 | 0 | 0
183.5 | 3.2256 | #DIV/0! | 0 | 0 | 0 | 0
184 | 3.2168 | #DIV/0! | 0 | 0 | 0 | 0
184.5 | 3.2081 | #DIV/0! | 0 | 0 | 0 | 0
185 | 3.1994 | #DIV/0! | 0 | 0 | 0 | 0
185.5 | 3.1908 | #DIV/0! | 0 | 0 | 0 | 0
186 | 3.1822 | #DIV/0! | 0 | 0 | 0 | 0
186.5 | 3.1737 | #DIV/0! | 0 | 0 | 0 | 0
187 | 3.1652 | #DIV/0! | 0 | 0 | 0 | 0
187.5 | 3.1568 | #DIV/0! | 0 | 0 | 0 | 0
188 | 3.1484 | #DIV/0! | 0 | 0 | 0 | 0
188.5 | 3.14 | #DIV/0! | 0 | 0 | 0 | 0
189 | 3.1317 | #DIV/0! | 0 | 0 | 0 | 0
189.5 | 3.1234 | #DIV/0! | 0 | 0 | 0 | 0
190 | 3.1152 | #DIV/0! | 0 | 0 | 0 | 0
190.5 | 3.107 | #DIV/0! | 0 | 0 | 0 | 0
191 | 3.0989 | #DIV/0! | 0 | 0 | 0 | 0
191.5 | 3.0908 | #DIV/0! | 0 | 0 | 0 | 0
192 | 3.0828 | #DIV/0! | 0 | 0 | 0 | 0
192.5 | 3.0748 | #DIV/0! | 0 | 0 | 0 | 0
193 | 3.0668 | #DIV/0! | 0 | 0 | 0 | 0
193.5 | 3.0589 | #DIV/0! | 0 | 0 | 0 | 0
194 | 3.051 | #DIV/0! | 0 | 0 | 0 | 0
194.5 | 3.0431 | #DIV/0! | 0 | 0 | 0 | 0
195 | 3.0353 | #DIV/0! | 0 | 0 | 0 | 0
195.5 | 3.0276 | #DIV/0! | 0 | 0 | 0 | 0
196 | 3.0199 | #DIV/0! | 0 | 0 | 0 | 0
196.5 | 3.0122 | #DIV/0! | 0 | 0 | 0 | 0
197 | 3.0045 | #DIV/0! | 0 | 0 | 0 | 0
197.5 | 2.9969 | #DIV/0! | 0 | 0 | 0 | 0
198 | 2.9894 | #DIV/0! | 0 | 0 | 0 | 0
198.5 | 2.9818 | #DIV/0! | 0 | 0 | 0 | 0
199 | 2.9743 | #DIV/0! | 0 | 0 | 0 | 0
199.5 | 2.9669 | #DIV/0! | 0 | 0 | 0 | 0
200 | 2.9595 | #DIV/0! | 0 | 0 | 0 | 0